## Slide 1: Clarity TRAINING
Installation & configuration

## Slide 2: content
Licensing & Software Protection
Installation
Registration
Station setup & configuration
INSTALLATION & CONFIGURATION / CONTENT

## Slide 3: Software protection
One HW key required for each station (PC)
One Clarity station running per PC
Non-driver USB HW key is standard (other types are obsolete)
INSTALLATION & CONFIGURATION / SOFTWARE PROTECTION

## Slide 4: Software protection
User code (16 position alphanumeric code) combined with the appropriate HW key (S/N yyyyyy) is required to operate Clarity (S/N xx-yyyyyy)
TRIAL option is available (HW key required) for potential customers' in‑house testing
Clarity rejects to start without a HW key plugged in
Clarity DEMO does not require a HW key
Important: Loss of the HW key is equal to a loss of the license!
INSTALLATION & CONFIGURATION / SOFTWARE PROTECTION

## Slide 5: Installation – checklist
Installation medium (USB flash or installer)
All required cables for devices ready
Local IT Administrator support (depending on company IT policies)
Obey rule: „Install drivers first!“ before connecting hardware (mainly for USB connection)
  Connect devices to PC after installation of Clarity is successfully finished.
Get familiar how to configure all devices in Clarity
  Check with respective manuals (if available) - chapter „3.x Clarity Configuration“
  Get familiar how to wire complex system for correct synchronization (mainly when using A/D converter)
INSTALLATION & CONFIGURATION / INSTALLATION CHECKLIST

## Slide 6: Installation – checklist
INSTALLATION & CONFIGURATION / INSTALLATION CHECKLIST
2
1
3
3
4
5
1. Hardware USB key
2. Installation USB with Clarity
3. Cables & wires (optional, order dependent)
4. A/D converter or another HW from DataApex portfolio
5. Documentation

## Slide 7: Installation
Clarity installer install.exe from USB or DataApex website (not applicable to OEM versions)
Use “Run as administrator“
Select language for installation
Confirm license agreement
Follow the guidance of the installer
Important: Do NOT connect any HW before installation (except for HW key)
INSTALLATION

## Slide 8: New installation
INSTALLATION / NEW INSTALLATION

## Slide 9: Types of installation
INSTALLATION / TYPES OF INSTALLATION
Typical installation is recommended
Some devices require additional parts – to be selected in the installation wizard e.g:
  Agilent HW (GC 8890…, Headspace AS, HPLC 1260/1290) → Instrument Control Framework (ICF)

## Slide 10: Update/Upgrade
INSTALLATION / UPDATE-UPGRADE
Choose Components and the rest as new installation

## Slide 11: MAJOR VERSION UPDATES SINCE VERSION 9.0
INSTALLATION / MAJOR VERSION UPDATES SINCE VERSION 9.0
Updates to major versions (change on the first digit place) are paid
  Validity of the entered user code is checked before installation

## Slide 12: A/D convertor connection
INSTALLATION / A/D CONVERTOR
Connect hardware only after finishing the software installation
Electric connections – the shield on signal cable must be connected (The shield on the analog input DET cable serves as common ground)
The digital input and outputs could be used for the start and ready synchronization

## Slide 13: First start - registration
Appears after the very first start
Required for updates/upgrades and support
Can be done later on the DataApex website (also possible from different PC)
CONFIGURATION / REGISTRATION

## Slide 14: General settings
Valid for the entire station
  User Accounts
  Configuration (& Units Setup)
  Directories
  GLP Options
  Digital Outputs
CONFIGURATION / GENERAL SETTINGS

## Slide 15: USER ACCOUNTS
CONFIGURATION / GENERAL SETTING / USER ACCOUNTS

## Slide 16: USER ACCOUNTS
CONFIGURATION / GENERAL SETTING / USER ACCOUNTS
Administration of user accounts
General settings of password rules
Setting of individual user privilages

## Slide 17: USER ACCOUNTS
CONFIGURATION / GENERAL SETTING / USER ACCOUNTS & USER DETAILS

## Slide 18: System COnfiguration
CONFIGURATION / GENERAL SETTING / SYSTEM CONFIGURATION

## Slide 19: System COnfiguration
CONFIGURATION / GENERAL SETTING / SYSTEM CONFIGURATION
Link between HW and Clarity
Left side:
  List of set control modules
  Overview of assignments to Instruments
Right side:
  Instrument Type
  Control modules “on“ Instrument
  Synchronization of HW with Clarity

## Slide 20: System COnfiguration
CONFIGURATION / GENERAL SETTING / SYSTEM CONFIGURATION
List of all available control modules
Search tool with filters
Available control modules

## Slide 21: System COnfiguration
CONFIGURATION / GENERAL SETTING / SYSTEM CONFIGURATION
Type enabled based on the User Code used
LC and GC by default
Switching ON the Extension
Instrument Type

## Slide 22: System COnfiguration
CONFIGURATION / GENERAL SETTING / SYSTEM CONFIGURATION
Synchronization of HW on Instrument with Clarity
Start Input – defines device that triggers analysis start
Ready Output – defines device that sets ready for next analysis (significant for Sequence)
Data inputs & outputs

## Slide 23: UNITs setup
CONFIGURATION / GENERAL SETTING / UNITS SETUP

## Slide 24: Units setup
CONFIGURATION / GENERAL SETTING / UNITS SETUP
Setting of used units
  Basic – Flow, Pressure....
  Auxiliary Signals – Voltage, Current...
  X-Axis (time) – minutes, seconds...
Set for each instrument separately
These setting are stored in the configuration *.cfg file

## Slide 25: Directories
CONFIGURATION / GENERAL SETTING / DIRECTORIES

## Slide 26: directories
Data and Audit Trail location
Directory = Where Clarity searches for Projects
Tool for networked use of Clarity with Clarity Offline
CONFIGURATION / GENERAL SETTING / DIRECTORIES

## Slide 27: Installation folder
CONFIGURATION / GENERAL SETTING / INSTALLATION FOLDER
Bin – the software itself, IQ, OQ
Cfg – configuration of the software
DataFiles – the chromatographic data
Documentation – Manuals, Datasheets
Uninstalll.exe – for software unistallation

## Slide 28: Installation folder
CONFIGURATION / GENERAL SETTING / INSTALLATION FOLDER
“Project_name“ folder – methods (*.met), seqeunces (*.seq)
  Calib subfolder – calibrations (*.cal), chromatograms of standards (*.prm)
  Data subfolder – sample chromatograms (*.prm)
Common – shared files = report styles, templates, etc.
Projects – individual project (*.prj) files
Datafiles

## Slide 29: Project setup
CONFIGURATION / GENERAL SETTING / PROJECT SETUP
Accessible from Instrument window
Administration of projects (opening, creation of new ones, renaming...)
Data and Calib folders of given project can be merged together in this dialog.

## Slide 30: Clarity2Go
CONFIGURATION / GENERAL SETTING / CLARITY2GO
Analysis monitoring from your smartphone or tablet
Available for Android and iOS

## Slide 31: GLP Options
CONFIGURATION / GENERAL SETTING / GLP OPTIONS

## Slide 32: GLP Options
CONFIGURATION / GENERAL SETTING / GLP OPTIONS
Necessary for work in a regulated environment
Turned off by default

## Slide 33: Digital outputs
CONFIGURATION / GENERAL SETTING / DIGITAL OUTPUTS

## Slide 34: Digital outputs
CONFIGURATION / GENERAL SETTING / DIGITAL OUTPUTS
Used to control the digital outputs of the installed A/D boards and other devices

## Slide 35: First start of instrument
CONFIGURATION / GENERAL SETTING / INSTRUMENT
Instrument window = center of analysis in Clarity
  Access to other windows + status information
Setup Configuration and other options as needed
Click Login
Select user name and project

## Slide 36: Thank you for your attention