## Slide 1: Clarity training
Advanced – part 2

## Slide 2: agenda
Chromatogram window
  Column setup, User columns, Noise & Drift
  Mathematical operations, Labels
  Result table, User variables
  Summary table, Performance table
  Force peak name, Peak coloring, Graph properties
Calibration
  LOD, LOQ
  Calibration Options, Manual Calibration
  Details of Calibration Point
  Bracketing and Standard Addition
Calculations
  ISTD, Normalized %
Groups
ADVANCED TRAINING

## Slide 3: agenda
Chromatogram window
  Column setup, User columns, Noise & Drift
  Mathematical operations, Labels
  Result table, User variables
  Summary table, Performance table
  Force peak name, Peak coloring, Graph properties
Calibration
  LOD, LOQ
  Calibration Options, Manual Calibration
  Details of Calibration Point
  Bracketing and Standard Addition
Calculations
  ISTD, Normalized %
Groups
ADVANCED TRAINING

## Slide 4: CHROMATOGRAM WINDOW
Comparing multiple chromatograms
  Max. number is set in User Options
Summary table shows results of the displayed chromatograms/signals
Hide/show signals via Ctrl+Click on signal color
OVERLAY
ADVANCED TRAINING / CHROMATOGRAM WINDOW / OVERLAY

## Slide 5: CHROMATOGRAM WINDOW
Mathematical operations between signals/chromatograms
Mathematical operations
ADVANCED TRAINING / CHROMATOGRAM WINDOW / MATHEMATICS

## Slide 6: CHROMATOGRAM WINDOW
Add text labels or lines
labels
ADVANCED TRAINING / CHROMATOGRAM WINDOW / LABELS

## Slide 7: Chromatogram window
Axis settings
Signal scaling –general settings, settings for individual signals
GRAPH PROPERTIES
ADVANCED TRAINING / CHROMATOGRAM WINDOW / GRAPH PROPERTIES

## Slide 8: Chromatogram window
Gradient display option
Auxiliary signals options
GRAPH PROPERTIES
ADVANCED TRAINING / CHROMATOGRAM WINDOW / GRAPH PROPERTIES

## Slide 9: CHROMATOGRAM WINDOW
Operable from the context menu of any table
Hide/Show columns
Change the order of columns
Setup columns
ADVANCED TRAINING / CHROMATOGRAM WINDOW / SETUP COLUMNS

## Slide 10: CHROMATOGRAM WINDOW
Access via context menu in Result/Summary table
Add custom calculations
Special values menu for more options
Specify user variables
User columns – RRT (Relative retention time)
ADVANCED TRAINING / CHROMATOGRAM WINDOW / USER COLUMNS
1
1

## Slide 11: CHROMATOGRAM WINDOW
Set noise interval
Add Signal/Noise calculation to a user column
User variables can be used
User columns – noise & drift
ADVANCED TRAINING / CHROMATOGRAM WINDOW / USER COLUMNS
1
1

## Slide 12: CHROMATOGRAM WINDOW
Manually assign response base and factor to unidentified peaks (peaks not present in calibration)
Calculation of amounts of unidentified (uncalibrated) compounds:
Amount = [Response Base] × [Response Factor]
Provides a “single-point quantification” of unidentified compounds (e.g., for universal response detectors like ELSD, FID, RI, etc.)
UNIDENTIFIED PEAKS – UNCALIBRATED RESPONSE
ADVANCED TRAINING / CHROMATOGRAM WINDOW / UNIDENTIFIED PEAKS – UNCALIBRATED RESPONSE

## Slide 13: CHROMATOGRAM WINDOW
Scale factor multiplies values in the Amount column and the Sample Amount (the entered units are used in the Amount when scaled, overriding the calibration settings)
  Method-specific (can be used only when calibration is set)
Dilution multiplies values in the Amount column
  Sample-specific (can also be entered via Sequence or Single Analysis)
SCALE & DILUTION
ADVANCED TRAINING / CHROMATOGRAM WINDOW / SCALE & DILUTION

## Slide 14: CHROMATOGRAM WINDOW
Analysis Variables are propagated from Single Analysis or Sequence to the measured chromatogram
  These can be used for custom calculations in the Result Table
User variables – Analysis variables
ADVANCED TRAINING / CHROMATOGRAM WINDOW / USER VARIABLES – ANALYSIS VARIABLES

## Slide 15: CHROMATOGRAM WINDOW
Analysis Variables are propagated from Single Analysis or Sequence to the measured chromatogram
  These can be used for custom calculations in the Result Table
User variables – Analysis variables
ADVANCED TRAINING / CHROMATOGRAM WINDOW / USER VARIABLES – ANALYSIS VARIABLES

## Slide 16: CHROMATOGRAM WINDOW
Method Variables are propagated from Method Setup – Advanced tab to the measured chromatogram
  These can be used for custom calculations in the Result Table
User variables – Method variables
ADVANCED TRAINING / CHROMATOGRAM WINDOW / USER VARIABLES – METHOD VARIABLES

## Slide 17: Chromatogram window
Results for all signals of selected chromatogram in one table
ALL SIGNALS RESULTS TABLE
ADVANCED TRAINING / CHROMATOGRAM WINDOW / ALL SIGNALS RESULT TABLE

## Slide 18: Chromatogram window
Summary of all opened chromatograms and signals based on calibration settings
Setup columns:
  Common for all compounds
  Compound-specific (Summary)
Summary table
ADVANCED TRAINING / CHROMATOGRAM WINDOW / SUMMARY TABLE

## Slide 19: Chromatogram window
Accessible through the context menu
Table inversion, header organization…
Summary table options
ADVANCED TRAINING / CHROMATOGRAM WINDOW / SUMMARY TABLE

## Slide 20: Chromatogram window
Set calculation type
Values for capacity ratio and efficiency from the used method
performance table
ADVANCED TRAINING / CHROMATOGRAM WINDOW / PERFORMANCE TABLE

## Slide 21: Chromatogram window
Set compound from calibration
Set custom name
Force peak name
ADVANCED TRAINING / CHROMATOGRAM WINDOW / FORCE PEAK NAME

## Slide 22: Chromatogram window
Set via Calibration,
Set peak area coloring via Graph Properties
Peak colors
ADVANCED TRAINING / CHROMATOGRAM WINDOW / PEAK COLORS

## Slide 23: agenda
Chromatogram window
  Column setup, User columns, Noise & Drift
  Mathematical operations, Labels
  Result table, User variables
  Summary table, Performance table
  Force peak name, Peak coloring, Graph properties
Calibration
  LOD, LOQ
  Calibration Options, Manual Calibration
  Details of Calibration Point
  Bracketing and Standard Addition
Calculations
  ISTD, Normalized %
Groups
ADVANCED TRAINING

## Slide 24: CALIBRATION
Limit of Detection (LOD) and Limit of Quantification (LOQ) – amounts lower than LOD/LOQ specified in calibration are marked accordingly in chromatogram Peak Type
LOD, LOQ
ADVANCED TRAINING / CALIBRATION / LOD, LOQ

## Slide 25: Calibration
Open via Calibration – Options or Calibration Options button
Display Mode – related to calculation types: ESTD, ISTD, NORM, STDADD
Calibration curve check – threshold used for the Failure message in Result Table – Peak Type
Automatic/Manual calibration
Calibration options
ADVANCED TRAINING / CALIBRATION / CALIBRATION OPTIONS

## Slide 26: Calibration
Accessible when adding points to calibration if Manual calibration is selected in Calibration Options
You can edit:
  Calibration level
  Identification criteria
  Compound information
  Response options (Area, Height, or Area %)
Manual response adjustment
Manual calibration
ADVANCED TRAINING / CALIBRATION / MANUAL CALIBRATION

## Slide 27: Calibration
When recalibrating a level using more points, you can view calibration point details for selected level (accessible also via Calibration – Show Details)
Checkbox state – include/omit from calibration
details OF Calibration point
ADVANCED TRAINING / CALIBRATION / DETAILS OF CALIBRATION POINT

## Slide 28: Calibration
Calibration in Method is used according to the Sequence Options – Calibration and sequence usage option
Name of calibration clone used if the method selected in Calibration and sequence usage supports cloning
BRACKETING AND STANDARD ADDITION
ADVANCED TRAINING / CALIBRATION / BRACKETING AND STANDARD ADDITION

## Slide 29: Calculations
Up to ten ISTD compounds can be present in the calibration/sample
2 types of ISTD calculations available:
  Any amount of ISTD (both calibration and sample chromatogram contain the used ISTD amount)
  Same amount of ISTD in samples and standards (ISTD amount = 0 in both locations)
ISTD
ADVANCED TRAINING / CALCULATIONS / ISTD

## Slide 30: Calculations
Amount % column contains the normalized percentage if all integrated peaks are calibrated
If the Sample Amount is 0, compounds in the Amount % column add up to a total of 100%
Otherwise, the value serves as the total sum of Amount %, and values in the Amount % column are normalized to the Amount set
NORM is the same as ESTD, except it additionally checks that both all identified peaks are calibrated, and all peaks from calibration are identified in the sample chromatogram
Normalized %
ADVANCED TRAINING / CALIBRATION / NORMALIZED %

## Slide 31: GROUPS
Set group ID and add peaks to the group
CREATING A GROUP
ADVANCED TRAINING / GROUPS / CREATING A GROUP
Add group in Calibration – set name and amount as for an ordinary peak

## Slide 32
Short break