## Slide 1: Clarity training
Advanced – part 3

## Slide 2: agenda
Batch processing
Report Setup
  Lab. header
  Tiled signals
  Order of sections
Export data
  Available options
Archive & Backup
Tips & Tricks
  Customization
  Post-run options
  Command Lock
  Shortcuts
ADVANCED TRAINING

## Slide 3: agenda
Batch processing
Report Setup
  Lab. header
  Tiled signals
  Order of sections
Export data
  Available options
Archive & Backup
Tips & Tricks
  Customization
  Post-run options
  Command Lock
  Shortcuts
ADVANCED TRAINING

## Slide 4: BATCH PROCESSING
BATCH
ADVANCED TRAINING / BATCH PROCESSING / BATCH
File Types available for Batch: Chromatograms, Calibration Standards, Sequence Files
Reprocessing chromatograms by selected method
Complete processing of whole sequences
Subsequent printing of results using a defined Report Style
Opening via a program of choice or further processing of the exported data

## Slide 5: Report setup
Window header contains the name of the current report style (depending on the window it was opened from)
Tabs which are to be printed are marked by a tick
Custom image/logo can be set in the Lab. Header of the report
Lab. header
ADVANCED TRAINING / REPORT SETUP / LAB. HEADER

## Slide 6: Report setup
Chromatogram – Tiled by
Tiled signals
ADVANCED TRAINING / REPORT SETUP / TILED SIGNALS

## Slide 7: Report setup
Order of all tabs between Lab. Header and Audit & Signatures can be changed by left-clicking the tab and dragging it to the desired position or by right-clicking and selecting where to move the selected tab
Order of sections
ADVANCED TRAINING / REPORT SETUP / ORDER OF SECTIONS

## Slide 8: agenda
Batch processing
Report Setup
  Lab. header
  Tiled signals
  Order of sections
Export data
  Available options
Archive & Backup
Tips & Tricks
  Customization
  Post-run options
  Command Lock
  Shortcuts
ADVANCED TRAINING

## Slide 9: Export data
Exporting data to other programs for further processing:
  Straightforwad way: Select, Copy & Paste (Ctrl+C, Ctrl+V)
  Automatic exports
Global settings applied in the Export Data dialog are common for:
  Manual export from the chromatogram window
  Automatic export via post run after each run end
  Automatic export via batch reprocessing or chromatograms/sequences
Export chromatograms is possible in various formats
OPTIONS
ADVANCED TRAINING / EXPORT DATA / OPTIONS

## Slide 10: Export data
Chromatogram Header = export basic chromatogram info (File Name, Sample ID, Sample, Instrument, Amount, Dilution)
Table Headers = column headers of the Result Table
Full format = filename, date, and time for each row of the Result Table
Append = export into a single file
Blank field in the File Name field indicates the use of the current chromatogram name + .txt
Instrument window
ADVANCED TRAINING / EXPORT DATA / INSTRUMENT WINDOW

## Slide 11: Export data
Additional export options are available; the Export Data dialog can be opened from Chromatogram as well
You can also select table content to Copy & Paste
Chromatogram window
ADVANCED TRAINING / EXPORT DATA / CHROMATOGRAM WINDOW

## Slide 12: Export data
When the Export Data option is enabled, the export is performed automatically post-run, according to the settings in the Export Data dialog
Chromatogram can be exported in different formats
SINGLE ANALYSIS POST RUN SETTINGS
ADVANCED TRAINING / EXPORT DATA / SINGLE ANALYSIS POST RUN SETTINGS

## Slide 13: Export data
Export data automatically after the chromatogram in the row is measured
The Export Data column can be added via Setup Columns
The export is performed according to the settings in the Export Data dialog
Together with the Export Data option, the Open option must be checked as well
SEQUENCE
ADVANCED TRAINING / EXPORT DATA / SEQUENCE

## Slide 14: Export data
Export data for multiple chromatograms using the Batch dialog
The export is performed according to the settings in the Export Data dialog
Chromatograms can be exported in various formats
Batch
ADVANCED TRAINING / EXPORT DATA / BATCH

## Slide 15: Export data
TEXT FILE FORMAT
ADVANCED TRAINING / EXPORT DATA / TEXT FILE FORMAT

## Slide 16: Export data
EXCEL (*.XLSX) FORMAT
ADVANCED TRAINING / EXPORT DATA / EXCEL (*.XLSX) FORMAT

## Slide 17: agenda
Batch processing
Report Setup
  Lab. header
  Tiled signals
  Order of sections
Export data
  Available options
Archive & Backup
Tips & Tricks
  Customization
  Post-run options
  Command Lock
  Shortcuts
ADVANCED TRAINING

## Slide 18: ARCHIVE & RESTORE
OVERVIEW
ADVANCED TRAINING / ARCHIVE & RESTORE / OVERVIEW
Data can be compressed to *.dgz archive files with reduced size
Writing into an existing *.dgz archive replaces all the preset data (appending is not possible)
Plan ahead: the project structure should allow you to archive whole projects when necessary
Access Create/Restore functions via Instrument
  Choose which file type to archive
Without compressing:
  When checked, no archive is created; the selected files are copied/moved
  When not checked, an archive containing selected files is created, located according to Target File Name

## Slide 19: TIPS & TRICKS
customization
ADVANCED TRAINING / TIPS & TRICKS / CUSTOMIZATION
You can customize Clarity to your preferred use in several ways
Open the Customize dialog by right-clicking the toolbar and selecting Customize…
  To add command icons to the toolbar, drag them from the menu in the Commands tab
  To change the Toolbar options, use the Toolbars tab or (un)select the corresponding items in the context menu

## Slide 20: TIPS & TRICKS
PERFORM POST RUN ACTIONS
ADVANCED TRAINING / TIPS & TRICKS / PERFORM POST RUN ACTIONS
PostRun Setting dialog is accessible from the menu File; the icon for Post Run can be added via Customize

## Slide 21: TIPS & TRICKS
COMMAND LOCK
ADVANCED TRAINING / TIPS & TRICKS / COMMAND LOCK
Pressing Ctrl when invoking a command will lock the command for repeated use
  This is useful when changing start/end for multiple peaks etc.
Terminate the command by Esc

## Slide 22: TIPS & TRICKS
Shortcuts
ADVANCED TRAINING / TIPS & TRICKS / SHORTCUTS
Most useful Clarity shortcuts

## Slide 23: Thank you for your attention
Next training coming soon