## Slide 1: validation

## Slide 2: VALIDATION LEVELS
VALIDATION / VALIDATION LEVELS
Checks version and integrity of Clarity files
IQ
Tests the overall analytical system performance
PQ
A/D Converter
Validator
OQ types:
Hardware (HW) generated data – Validator
Software (SW) generated data – Clarity Virtual Detector
Tests data acquisition and processing
OQ
Installation
Qualification
Operational
Qualification
Performance
Qualification

## Slide 3: INSTALLATION QUALIFICATION (IQ)
Checks version and integrity of Clarity files
Run from main Clarity window – Help menu
VALIDATION / IQ

## Slide 4: Iq report
VALIDATION / IQ REPORT
General infromation
Overall IQ status
  Passed
  Passed with warnings
  Failed
Status of each file

## Slide 5: Iq reporT
VALIDATION / IQ REPORT
Run separately
Must be enclosed with Clarity IQ
3rd partY package

## Slide 6: operational QUALIFICATION (oQ)
Tests data acquisition and processing
Tool supporting 21 CFR Part 11 Regulations
Requires System Suitability Test (SST) extension
Validation kit required only for stations with A/D converters
VALIDATION / OQ

## Slide 7: OQ
Validator
AC/DC power adapter
Cable for A/D converter
Terminal board
Validator certificate
validation kit
VALIDATION / OQ / VALIDATION KIT

## Slide 8: OQ
Validator
  Station with A/D converter
Virtual detector
  Station with only digital acquisition
Generating chromatograms
VALIDATION / OQ / GENERATING CHROMATOGRAMS
A/D Converter
SST
+
Validator
SST
Virtual Detector
+

## Slide 9: OQ
Run from …\Clarity\Bin\OQ_VALIDATION\
Avoid printers requiring user response
Wizard 1/5
VALIDATION / OQ / WIZARD

## Slide 10: OQ
Select validation type
  Do you use A/D converter?
Wizard 2/5
VALIDATION / OQ / WIZARD

## Slide 11: OQ
Exclusive to A/D converter validation
  Select your HW
  Let the Validator stabilize
  Change Validator data set
Wizard 3/5
VALIDATION / OQ / WIZARD

## Slide 12: OQ
Set project name
  All OQ data will be stored into it
Check the summary
Wizard 4/5
VALIDATION / OQ / WIZARD

## Slide 13: OQ
Login to Clarity
For A/D converter validation
  Press the button on Validator
Virtual detector validation starts automatically
Wizard 5/5
VALIDATION / OQ / WIZARD

## Slide 14: OQ
3 sets of analyses within 1 sequence
  Calibration curve linearity
  Repeatability and precision
  Acquisition range
Each set associated with different method and report style
sequence
VALIDATION / OQ / SEQUENCE
Linearity
Repeatability
Range

## Slide 15: OQ
Project folder named according to OQ wizard
2 PDF in Data, 1 PDF in Calib
Results 1/3 – DATA
VALIDATION / OQ / RESULTS

## Slide 16: OQ
1st set must be checked manually
2nd and 3rd are checked by SST
Results 2/3 - Evaluation
VALIDATION / OQ / RESULTS

## Slide 17: OQ
3 PDF files from OQ project folder
Protocol
Results 3/3 – OQ report
VALIDATION / OQ / RESULTS

## Slide 18: OQ
Personnel trained in validation
Certified by DataApex
Certification valid for 5 years
Who can perform?
VALIDATION / OQ / WHO CAN PERFORM

## Slide 19: OQ
Validation kit manual
  Tested parameters
  SST limits
  Validator chromatograms
Additional info
VALIDATION / OQ / ADDITIONAL INFO

## Slide 20: Performance qualification (pq)
Compliance with SOPs and workflow of given workplace
SST extension can be used
Must be prepared by the workplace personnel based on:
  Used instruments
  SOPs
  Applicable regulations
IQ and OQ must be successful before doing PQ
No certificate for PQ is issued by DataApex, , but support is available if required
VALIDATION / PQ

## Slide 21: Thank you for your attention